## front page
Kwiktrak final installation checklist | Date:______________________
Facility:____________________________________________________ | Inspector Name:____________________ |  |  | Signature:______________________
SO #: ______________________________________________________ |  | (produce KwikSketch and notes on page 2 if layout is different than signed drawing)
(correct = O, discrepancies = X + describe on page 2) room numbers
overall structure description:
track respects transfer zones (draw room layout on page 2 if different)
respect minimum sprinkler distances (see Kwiktips)
installation avoids obstructions (lights, vents, curtains…)
respect minimum distances for lift insertion (11 inches or 18 inches)
brackets are in locked position
alignment of the tracks respects straight portion, smooth lines
track junction showing gaps smaller than 1mm (1/16 inch)
absence of abusive accumulation of shims (less 3 shims each size / bracket)
end stopper in place, attached, set screw tightened & adjusted
maximum rail cantilever is 12 inches
WLT certification sticker properly installed and signed
minimum damages to the ceiling, walls and tracks
transformer box properly installed, charging correctly
excess charger wires hidden in plastic conduit, clean installation
access doors properly installed (see Kwiktips)
anchors / hardware respects methods and specifications (see Kwiktips)
respect technical shop drawings for this particular structure (see Kwiktips)
sandwich effect applied correctly, proper stiffness of the tracks
lateral reinforcements installed properly when / where required (see Kwiktips)
all parts of the installation are locked, tightened and solidly attached
any material left on site: no yes |  |  | pack up material, fill merchandise return form and call coordinator
KwikSketch and notes (room layout, bed / toilet / tub position, sprinkler location, charger position, discrepancies)
room | room
room | room
room | room
room | room
Room #:
Room #:
Room #:
Room #:
Room #:
Room #:
Room #:
Room #: